## Slide 1: ERGAssist User Guide
Please Cite Feola et al. “Development of An Automated Electroretinography Analysis Approach” when using ERGAssist

## Slide 2: Index
User Guide
  Installation….……………………………………………………
  GUI Overview.….………………………………………………
  Use.….………………………………………………………………
  Outputs ……………………………………………………………
.

## Slide 3: Installation
Download and Install Matlab MyAppInstaller_mcr Package from web (Check Readme file)
Download ERGAssist App
Run ERGAssist

## Slide 4: GUI Overview

## Slide 5: General Information
Time – units ms
Voltage – microvolts (except Celeris where voltage is nanovolts [nV])
Flash Stimuli – log cd*s/m2 (except RetEval)
Users can load multiple files that have the same format and steps at the same time
For Toggle Buttons: Orange highlighted = ON

## Slide 6: ERGAssist Layout

## Slide 7: ERG Assist
Load Waveforms
  Allows users to upload waveforms
Analyze ERG
  Analyzes ERGs
Continue Session
  Allows users to load an exported session (from ‘Save Session’) or Automatically exported on export
  File type is ‘.mat’
Export Results
  Exports marked ERG data into Excel
  Also Exports session if ‘Auto Save’ is highlighted

## Slide 8: Protocol Information
Scotopic Steps
  Allows users to input number of dark-adapted steps (called scotopic here)
Photopic Steos
  Allows users to input number of light-adapted steps (called photopic here)
Flicker
  Toggle ON / OFF if there was a Flicker step
Update Steps
  Updates Table to reflect steps
Load Flash Int
  Allows user to upload flash information (Examples 2 and 3)
Load Time Data
  Toggle ON/OFF if Excel data includes time information
Flicker Frequency
  Allows user to input flicker sampling frequency (only used if Time data is not included)
Preflash Stimulus
  Time prior to the flash stimuli (only used if time data is not included)
ERG System
  Toggle between how the data was collected/exported

## Slide 9: Processing Options
Species
  Allows users to adjust between rat/mouse/human inputs
Sampling Freq
  Allows users to input sampling frequency to generate time data (only used if Time data is not included)
OPFilter Type
  Allows user to toggle between Butterworth and FIR Digital Filter
OP Filter Order
  Allows user to adjust the order of the Butterworth filter
Low Pass
  Allows user to adjust the low pass filter frequency (Hz)
  Used in low-pass and bandpass filter (OP)
High Pass
  Allows user to adjust the high pass filter frequency (Hz)
  Used in bandpass filter (OP)
  Note: Data will need to be reanalyzed to see changes in filters

## Slide 10: Processing Options
Intensity Independent
  Allows users to adjust to ignoring flash stimuli information upon a reanalysis if desired
Filter Noisy Curves
  Toggles ON / OFF automatically looking for noisy curves based on frequencies in the signals
Save Session
  Allows user to pause analysis and save data as a ‘.mat’ file.
  This can be loaded to complete analysis or inspection later
Auto Save
  When pressed, upon export a session (‘.mat’ file) will automatically be saved as well with the filename on exported with the addition tag '_Autosave'

## Slide 11: Edit PII & PIII Marks
Adjust a-wave
  Moves or adds a-wave mark (red square)
Remove a-wave
  Removes a-wave mark
Adjust b-wave
  Moves or adds a-wave mark (red circle)
Remove b-wave
  Removes b-wave mark

## Slide 12: Adjust STR & PhT Marks
Pos. Threshold
  Moves or adds positive STR or c-wave peak (blue diamond)
Remove Pos.
  Removes positive STR or c-wave peak (blue diamond)
Neg. Threshold
  Moves or adds negative STR or negative photopic threshold (blue diamond)
Remove Neg.
  Removes negativeSTR or c-wave peak (blue diamond)

## Slide 13: Edit OP Marks
Add OP
  Add OP (start with trough then peak)
Remove OP
  Removes selected OP (select trough then peak)
Assist Auto OPs
  Clears OPs and allows user to select trough of first OP. Others are automatically selected from this point
Clear OPs
  Removes all OPs
  Note: ERGAssist currently only allows 5 OPs to be marked

## Slide 14: Display Options
Raw Waveform
  Displays the raw ERG data (black line) underneath the processed waveform
Vertical Lines
  Displays dashed vertical lines at the location of the a-wave and b-wave
OPs
  Toggles ON/OFF oscillatory potential waveform from bandpass filter

## Slide 15: Figure Window
Allows adjustment of time displayed
Start
  Beginning of x-axis display
End
  End of x-axis display
Use Figure Window
  Toggles ON / OFF using Figure Window range on the plots

## Slide 16: Blind Waveforms
Toggles ON /OFF blinding of ID information for blinded analysis

## Slide 17: Misc. Items
Clear Markings
  Clears all markings from current waveform
Adjust Baseline
  To adjust the reference point (baseline) that is considered 0 Volts
Export Waveform
  Exports current waveform data (voltage, time, and OP data)

## Slide 18: Example 1 Loading From Excel
Manually Input Flash and Time Information

## Slide 19: Key Information For Example 1
User Needs to Update:
  Flash Type: Dark Adapted (Scotopic), Light Adapted (Photopic), Flicker
    User also inputs flash stimuli information
  Sampling Frequency
  Flicker Frequency (if applicable)
  Pre-stimulus time
Generates time data from User Input
  Time waveform is generated from sampling frequency (Hz)
  No time data is included in the Excel file (only voltage response)
  Voltage data is in microvolts

## Slide 20
Shows:
Loading step type and manually inputting flash stimuli (flash intensities)
Illustrating toggling on/off OP waveforms

## Slide 21: Example 2 Loading From Excel
Load Input Flash Stimuli and Time Information

## Slide 22: Key Information For Example 2
User Needs to Update:
  Flash Steps: Dark Adapted (Scotopic), Light Adapted (Photopic), Flicker
  Load file containing flash stimuli information
    Can be a single row or column
    File format: txt, xlsx, xls, or csv
Loads Time Data from User Input
  Flash log cd-s/m2
  Excel file contains time (ms) and voltage (microV) information

## Slide 23
Shows:
Loading flash stimuli (flash intensities) after determining step type
Illustrating automatically flagged waveforms and removing a flag

## Slide 24: Example 3 Loading From Excel
Load Input Flash and Flash Type and Time Information

## Slide 25: Key Information For Example 3
User Needs to Update:
  Flash Steps: Dark Adapted (Scotopic), Light Adapted (Photopic), Flicker
  Load file containing flash stimuli information
    Should be two rows
    File format: txt, xlsx, xls, or csv
    Step Key
      0 – Dark Adapted (scotopic); 1 – Light Adapted (photopic); 2 – Flicker
Loads Data from User Input
  Excel file containing voltage (microV) information

## Slide 26
Shows:
Loading flash stimuli (flash intensities) and type of flash
Changing species from rat to mouse
Noisy curve filtering
  Results in flagged waveforms (black highlight)
Renanalzying data without filtering out noisy curve
  Results in flagged waveforms (blue highlights)
  Should be checked
Shows Adjusting placement of b-wave amplitude

## Slide 27: Example 4 Loading From LKC Device

## Slide 28: Key Information For Example 4
User Needs to Update:
  Flash Steps: Dark Adapted (Scotopic), Light Adapted (Photopic), Flicker
  Directly load LKC data from Bigshot or SunBurst
Loads Data from User Input
  Directly load LKC data from
    Time should be in ms and Voltage should be in microV
Note: This Example shows using ‘Estimate Step Type’ Import (beta Version)

## Slide 29
Shows:
When you do not adjust flash stimuli
Highlights feature to have the program estimate the type of waveform (dark-adapted, light-adapted, or flicker)
  Note: This is a beta version and classifications were trained on rat data collected on an LKC device.
Shows using ‘Assist Auto OPs’
  Clears all OPs and allows the user to manually mark the first trough (the program will then automatically mark OP1-5 based on this placement

## Slide 30: Example 5 Loading From Celeris

## Slide 31: Key Information For Example 4
User Needs to Update:
  Flash Steps: Dark Adapted (Scotopic), Light Adapted (Photopic), Flicker
  Directly load Celeris file
Loads Data from User Input
  Directly load LKC data from
    Time should be in ms and Voltage should be in nanoV
Note: This Example shows using ‘Continue Without Flash Intensity’ Import

## Slide 32
Shows:
When you do not adjust flash stimuli
Highlights feature to have the program ignore conditions and analyze independently.
  Note: A flicker waveform will be analyzed the same including looking for a-wave, b-wave and OPs. These will need to be reanalyzed manually.
For Celeris data (if you have all dark-adapted steps) you can import the flash information directly
Shows using ‘Clear OPs’
  Clears all OPs
Shows ‘Adjust a-wave’ and ‘Adjust b-wave’ butons
  This allows you to manually adjust markings

## Slide 33: Output Format

## Slide 34: Excel Output
Subject ID
Eye
ERG
Test Date
Step
Number
Flag (and note)
Output Parameter
DATA
DATA
DATA

## Slide 35: Output Parameters:
A Amp – a-wave amplitude
A IT – a-wave implicit time
B Amp – b-wave amplitude
B IT – b-wave implicit time
OP1 IT – oscillatory potential 1 implicit time
OP1 Amp – oscillatory potential 1 amplitude
OP information is repeated for OP 1 -5
pSTR Amp – positive scotopic threshold (STR) amplitude (also c-wave)
pSTR IT – positive STR implicit time (also c-wave)
nSTR Amp – negative STR amplitude (also negative photopic threshold response)
nSTR IT – negative STR implicit time (also negative photopic threshold response)
Note: all exported amplitudes are in microV (including Celeris) and implicit times are in ms

## Slide 36: ERG Waveform File Formats
Generic
  Excel
LKC
  Bigshot
  Sunburst
  RetEval
Diagnosys
  Celeris

## Slide 37: General Data Organization
Dark Adapted Steps are First
Dark Adapted Steps are Referred to as Scotopic (Sc)
Light Adapted Steps are Referred to as Photopic (Ph)
Flicker
  Flicker is the last step
  Only a single Flicker Step per eye
Note: This does not apply if you are Ignoring Flash Information or Estimating Step Type (akin to Examples 4 and 5)

## Slide 38: CHECK DATA
ERGAssist helps in quickly analyzing ERG data, but it is STRONGLY encouraged that all waveforms are quickly examine each waveform to ensure markers are accurate and align with your laboratories practices in marking waveforms
If there is a bug or desired change. Please communicate with the Corresponding Author of “Development of An Automated Electroretinography Analysis Approach” Andrew Feola, Ph.D. @ andrew.feola@emory.edu